# NIYAMA_3DP  RS485 / UART Protocol

Segment-LCD environmental monitor, MM32F0141C6P.   Firmware 1.2.2.

This document is generated from the firmware sources. Every parameter listed was taken from the read and write dispatch tables in main.c, so it describes what the code does rather than what it was once intended to do.

# 1.  Link settings

| Setting | Value |
| --- | --- |
| Signalling | RS485 half duplex, or plain UART |
| Baud rate | 57600 by default; settable to 9600 or faster (see 0x41) |
| Frame | 8 data bits, no parity, 1 stop bit, no flow control |
| Device address | 1 by default, 0 addresses every device at once |
| Byte order | Binary fields are little endian |
| Epoch | Seconds since 1 January 1970, in the device’s own local time |

# 2.  Frame format

A request:

FF   ID   CMD   PID   [ payload ... ]   CRC   FE
  0    1    2     3     4 ...             n-2   n-1

A reply:

FD   ID   CMD   STATUS   PID   [ payload ... ]   CRC   FC
  0    1    2     3        4     5 ...             n-2   n-1

Fields

| Byte | Meaning |
| --- | --- |
| FF / FD | Start of a request / start of a reply |
| ID | Device address. The reply always carries the device’s own address, not the one that was asked for. |
| CMD | 0x10 read, 0x11 write, 0x00 read-with-alarm |
| STATUS | Reply only. See section 4. |
| PID | Parameter identifier. The reply echoes the one from the request — check it, see section 11. |
| CRC | See section 3 |
| FE / FC | End of a request / end of a reply |

Read-with-alarm carries no parameter identifier

0x00 returns the live readings together with the alarm state, so there is nothing to name and the PID byte is absent. Its request is five bytes, not six:

FF   ID   00   CRC   FE
  0    1    2    3     4

  device 1:   FF 01 00 56 FE      0x55 + 0x01 + 0x00 = 0x56

Five is the shortest frame the device accepts. Anything shorter is discarded without a reply: the checksum length is derived from the frame length, and below five bytes there is nothing left for it to cover.

Addressing by serial number

A request may be addressed to a serial number instead of a device address. It opens with 0xEA followed by the last 8 characters of the serial number, and closes with 0xEB in place of 0xFE. Only the device whose serial number matches will answer.

# 3.  Checksum

An additive checksum folded into seven bits. It covers every byte from index 1 up to but not including the checksum itself.

total = 0x55;
for (i = 1; i < crc_index; i++) total += frame[i];
if (total > 0x7F) total &= 0x7F;

Worked example — read the firmware version from device 1:

request   FF 01 10 34 1A FE
            0x55 + 0x01 + 0x10 + 0x34 = 0x9A, over 0x7F, so & 0x7F = 0x1A

# 4.  Status byte

A bit field. Zero means everything is well. More than one bit can be set at once.

| Bit | Name | Meaning |
| --- | --- | --- |
| 0x02 | INVALID_PARA | The device does not support that parameter in this build |
| 0x04 | DP1_FAULTY | Differential pressure sensor 1 not responding |
| 0x08 | RH_TEMP_FAULTY | Temperature / humidity sensor not responding |
| 0x10 | DP2_FAULTY | Differential pressure sensor 2 not responding |
| 0x20 | DP3_FAULTY | Differential pressure sensor 3 not responding |
| 0x40 | RTC_INVALID | The clock has lost integrity; timestamps are not trustworthy |

A reply to an unsupported parameter still arrives, with INVALID_PARA set. There is one exception, described in section 11.

# 5.  Reading and writing a value

Read

request   FF  ID  10  PID  CRC  FE
  reply     FD  ID  10  ST  PID  <ASCII digits>  CRC  FC

The value comes back as ASCII decimal with no padding, and a leading minus sign when it is negative. Its length therefore varies.

Write

request   FF  ID  11  PID  <ASCII digits>  CRC  FE
  reply     FD  ID  11  ST  PID  <the same digits>  CRC  FC

The reply echoes the request from byte 3 onward. It confirms the frame was accepted, not that the value was stored — a value outside the parameter’s permitted range is discarded silently. Read the parameter back to be sure.

Value range

Every ASCII value is parsed into a signed 16 bit integer. Anything outside −32768 to 32767 wraps and is stored as something else entirely, so a host must not send it.

# 6.  Scaling

Scaling is not uniform. Each parameter’s factor is given in the tables in section 8; the groups are:

| Group | On the wire | Example |
| --- | --- | --- |
| Alarm setpoints | Hundredths of the unit (stored as tenths) | 5500 = 55.00 Pa, stored as 550 |
| Recorded minimum and maximum | Hundredths of the unit | 2524 = 25.24 °C |
| DP reading clamp (0x6E) | Tenths of a Pa | 3000 = 300.0 Pa |
| Per-slot DP offset (0x72) | Tenths of a Pa | −0025 = −2.5 Pa |
| DP zero offset (0x71) | Hundredths of a Pa | +00125 = 1.25 Pa |
| DP span factor (0x5F) | Hundredths of a Pa | −00250 = −2.50 Pa |
| Temperature / humidity calibration | Tenths of the unit | 00250 = 25.0 |
| Everything else | The plain number | 10 = 10, in whatever unit the parameter column gives |

# 7.  Timing

The device serves one message at a time and cannot take in the next while it is replying. A host should leave roughly 80 ms between transactions.

If a reply does not arrive, let the line go quiet for a few hundred milliseconds before asking again; that is what clears the device’s receiver.

Firmware before 1.0.4 was considerably less forgiving: a failed checksum latched its receive-pending flag and gated the receive interrupt permanently, so the unit went deaf for seconds at a time. Back-to-back frames lost 43 of 60 on the bench. In 1.0.4 the same test loses 2 of 60.

# 8.  Parameters

Access is what the firmware actually implements: R where only the read dispatch handles the identifier, W where only the write dispatch does, RW where both do. Build says which firmware build the parameter exists in — asking a DP1+DP2+DP3 unit for a humidity parameter returns INVALID_PARA, and the other way round.

## 8.1  Alarm setpoints

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x02 | DP1 upper alarm ON | Pa | RW | both | hundredths |
| 0x01 | DP1 upper alarm OFF | Pa | RW | both | hundredths |
| 0x03 | DP1 lower alarm OFF | Pa | RW | both | hundredths |
| 0x04 | DP1 lower alarm ON | Pa | RW | both | hundredths |
| 0x06 | DP2 upper alarm ON | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x05 | DP2 upper alarm OFF | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x07 | DP2 lower alarm OFF | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x08 | DP2 lower alarm ON | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x62 | DP3 upper alarm ON | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x61 | DP3 upper alarm OFF | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x63 | DP3 lower alarm OFF | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x64 | DP3 lower alarm ON | Pa | RW | DP1+DP2+DP3 | hundredths |
| 0x0A | Temperature upper alarm ON | deg | RW | DP1+Temp+RH | hundredths |
| 0x09 | Temperature upper alarm OFF | deg | RW | DP1+Temp+RH | hundredths |
| 0x0B | Temperature lower alarm OFF | deg | RW | DP1+Temp+RH | hundredths |
| 0x0C | Temperature lower alarm ON | deg | RW | DP1+Temp+RH | hundredths |
| 0x0E | Humidity upper alarm ON | %RH | RW | DP1+Temp+RH | hundredths |
| 0x0D | Humidity upper alarm OFF | %RH | RW | DP1+Temp+RH | hundredths |
| 0x0F | Humidity lower alarm OFF | %RH | RW | DP1+Temp+RH | hundredths |
| 0x10 | Humidity lower alarm ON | %RH | RW | DP1+Temp+RH | hundredths |

## 8.2  Recorded extremes

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x1C | DP1 recorded minimum | Pa | R | both | hundredths |
| 0x1D | DP1 recorded maximum | Pa | R | both | hundredths |
| 0x1E | DP2 recorded minimum | Pa | R | DP1+DP2+DP3 | hundredths |
| 0x1F | DP2 recorded maximum | Pa | R | DP1+DP2+DP3 | hundredths |
| 0x65 | DP3 recorded minimum | Pa | R | DP1+DP2+DP3 | hundredths |
| 0x66 | DP3 recorded maximum | Pa | R | DP1+DP2+DP3 | hundredths |
| 0x20 | Temperature recorded minimum | deg | R | DP1+Temp+RH | hundredths; converted to Fahrenheit when 0x2E is 1 |
| 0x21 | Temperature recorded maximum | deg | R | DP1+Temp+RH | hundredths; converted to Fahrenheit when 0x2E is 1 |
| 0x22 | Humidity recorded minimum | %RH | R | DP1+Temp+RH | hundredths |
| 0x23 | Humidity recorded maximum | %RH | R | DP1+Temp+RH | hundredths |

## 8.3  Clock

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x11 | Clock, hour | h | R | both | plain |
| 0x12 | Clock, minute | min | R | both | plain |
| 0x13 | Clock, second | s | R | both | plain |
| 0x16 | Clock, day of month |  | R | both | plain |
| 0x15 | Clock, month |  | R | both | plain |
| 0x14 | Clock, year |  | R | both | plain, four digits |
| 0x4B | Date and time in one exchange |  | RW | both | see section 10 |

## 8.4  Device and identity

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x34 | Firmware version |  | R | both | major*100 + minor*10 + patch, so 122 is 1.2.2 |
| 0x1A | Device address |  | RW | both | plain; must be written twice, see section 10 |
| 0x40 | Serial number |  | RW | both | 16 ASCII characters, see section 10 |
| 0x41 | Baud rate code |  | RW | both | see section 9; only 3 to 9 are accepted |
| 0x2E | Temperature unit |  | RW | DP1+Temp+RH | 0 Celsius, 1 Fahrenheit |
| 0x19 | Log interval | min | RW | both | plain, 1 to 1440; see section 11 |
| 0x58 | Master enable |  | RW | both | plain |
| 0x54 | 24 hour mean start hour | h | R | both | plain |
| 0x4E | Set the displayed-parameter word |  | RW | both | password then bit pattern, see section 10 |

## 8.5  Alarm behaviour

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x35 | Buzzer ON time | s | RW | both | plain; 0 in either time silences it |
| 0x36 | Buzzer OFF time | s | RW | both | plain; 0 in either time silences it |
| 0x3A | Alarm acknowledge silence time | s | RW | both | plain |
| 0x3C | Acknowledge the current alarm |  | W | both | value at offset 5, not 4 |
| 0x3B | Acknowledge password |  | RW | both | value at offset 5, not 4 |
| 0x60 | Relay control |  | W | both | relay digit, then 0 or 1 |

## 8.6  Sensing times

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x5C | DP1 alarm sensing time | s | RW | both | plain |
| 0x5D | DP2 alarm sensing time | s | RW | DP1+DP2+DP3 | plain |
| 0x68 | DP3 alarm sensing time | s | RW | DP1+DP2+DP3 | plain |
| 0x5B | Door sensing time | s | RW | both | plain |
| 0x5A | Door contact polarity |  | RW | both | plain |

## 8.7  Display and communications

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x5E | LCD brightness |  | RW | both | plain, 0 to 15 |
| 0x6F | LCD disable |  | RW | both | 1 switches the display off |
| 0x70 | UART disable |  | RW | both | 1 switches the port off; see section 11 |
| 0x69 | Unprompted send interval | s | RW | both | plain, 0 disables it |
| 0x4F | Parameter scroll time | s | RW | both | plain, 1 to 60; ONE-LINE panels only - a three-line build answers INVALID_PARA |

## 8.8  Calibration

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x37 | Unlock factory calibration |  | W | both | send the factory password |
| 0x38 | Unlock customer calibration |  | W | both | send the customer password |
| 0x39 | Customer password |  | RW | both | plain, up to 999 |
| 0x4D | Factory customer password |  | RW | both | plain, up to 9999 |
| 0x30 | DP1 calibration | Pa | RW | both | tenths, see section 10 |
| 0x31 | DP2 calibration | Pa | RW | DP1+DP2+DP3 | tenths, see section 10 |
| 0x67 | DP3 calibration | Pa | RW | DP1+DP2+DP3 | tenths, see section 10 |
| 0x32 | Temperature calibration | deg | RW | DP1+Temp+RH | tenths, see section 10 |
| 0x33 | Humidity calibration | %RH | RW | DP1+Temp+RH | tenths, see section 10 |
| 0x71 | DP zero offset | Pa | RW | both | hundredths, indexed, see section 10 |
| 0x5F | DP span factor | Pa | RW | both | hundredths, indexed, customer password only |
| 0x6E | DP reading clamp | Pa | RW | both | tenths, indexed, see section 10 |
| 0x72 | DP per-slot offset | Pa | RW | both | tenths, indexed, see section 10 |
| 0x52 | Reset calibration |  | W | both | channel digit 0, 1 or 2 |

## 8.9  Logs and bulk reads

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x4A | Records in the regular log |  | R | both | the write index, which is the count until the ring wraps |
| 0x49 | Read the log from a date |  | R | both | binary block |
| 0x45 | Recent readings, all of them |  | R | both | binary block, 1507 bytes |
| 0x46 | Recent readings, one index |  | R | both | binary block |
| 0x47 | 24 hour log, one index |  | R | both | binary block |
| 0x4C | 24 hour log, current index |  | R | both | plain |
| 0x53 | Min/max/mean log, one index |  | R | both | binary block |
| 0x48 | Live values |  | R | both | binary, see section 10 |

## 8.10  Maintenance

| ID | Parameter | Unit | Access | Build | Notes |
| --- | --- | --- | --- | --- | --- |
| 0x50 | Erase the external flash |  | W | both | no value |
| 0x51 | Reset the clock to its default |  | W | both | no value |
| 0x3F | Broadcasting on or off |  | RW | both | write a duration in minutes to start it; a read answers 0 or 1 |
| 0x3E | Stop broadcasting |  | W | both | no value |

# 9.  Baud rate codes

Parameter 0x41 carries a code, not a rate. The firmware only accepts 3 to 9; codes 0 to 2 exist in the table but are rejected on write, and a stored value outside 3 to 9 is replaced at boot.

| Code | Rate | Accepted |
| --- | --- | --- |
| 0 | 1200 | no |
| 1 | 2400 | no |
| 2 | 4800 | no |
| 3 | 9600 | yes |
| 4 | 14400 | yes |
| 5 | 19200 | yes |
| 6 | 28800 | yes |
| 7 | 38400 | yes |
| 8 | 57600 | yes, the default |
| 9 | 115200 | yes |

The device switches rate as soon as it has sent the reply, so the host must reconnect at the new rate.

# 10.  Frames that are not plain numbers

## 10.1  Date and time in one exchange  (0x4B)

Read

request   FF  ID  10  4B  CRC  FE
  reply     FD  ID  10  ST  4B  DDMMYYhhmmss  CRC  FC       (19 bytes)

Twelve ASCII digits: day, month, two-digit year, hour, minute, second.

Write

request   FF  ID  11  4B  DDMMYYhhmmss  CRC  FE

The device refuses a time earlier than the one it already holds, so a clock that is running fast cannot be wound back this way.

The field carries a two digit year, so only 2000 to 2099 can be expressed. The conversion to epoch seconds is undefined below 2000 - the year is held unsigned and would wrap - but no such date can be entered through this field in the first place.

The epoch the device reports is seconds since 1 January 1970 measured against its own local time. No timezone or daylight-saving correction is applied anywhere, so a host comparing device timestamps with a UTC source has to account for the offset itself.

## 10.2  Serial number  (0x40)

Read

request   FF  ID  10  40  CRC  FE
  reply     FD  ID  10  ST  40  <16 ASCII characters>  CRC  FC     (23 bytes)

Write

request   FF  ID  11  40  <16 ASCII characters>  CRC  FE

Exactly sixteen characters, no length field and no padding. They are free-form, so one of them could be 0xFC; read the reply by its fixed length of 23 bytes rather than hunting for the terminator.

The last eight characters are what serial-number addressing matches against (section 2).

## 10.3  Live values  (0x48)

The only reply that is binary rather than ASCII. Fixed at 51 bytes.

request   FF  ID  10  48  CRC  FE

  reply, 51 bytes:
    [0]      FD
    [1]      device address
    [2]      command
    [3]      status
    [4]      48
    [5..8]   epoch seconds            uint32, little endian
    [9]      sensor fault flags       same bits as the status byte
    [10..13] DP1                      float
    [14..17] DP2   or temperature     float
    [18..21] DP3   or humidity        float
    [22..25] DP1 minimum              float
    [26..29] DP1 maximum              float
    [30..33] DP2 / temperature min    float
    [34..37] DP2 / temperature max    float
    [38..41] DP3 / humidity min       float
    [42..45] DP3 / humidity max       float
    [46]     DP1 alarm                0 none, 1 high, 2 low
    [47]     DP2 / temperature alarm
    [48]     DP3 / humidity alarm
    [49]     CRC over [1..48]
    [50]     FC

Read this one by length. The floats are raw IEEE-754 and any of their bytes can be 0xFC, so scanning for the terminator will truncate the frame sooner or later.

Nothing in the frame says which of the two channel layouts applies. Ask for 0x2E, which only exists in the DP1+Temp+RH build: an answer means temperature and humidity, INVALID_PARA means DP2 and DP3.

## 10.4  Readings with alarms  (command 0x00)

One exchange for everything the front panel shows. The request is five bytes and names no parameter — see section 2. The reply is variable length, 43 bytes on a DP1 + temperature + humidity build:

FD  ID  00  STATUS    <reading 1>  EE  <reading 2>  EE  <reading 3>  EE    <alarm 1><alarm 2><alarm 3>  EE    <door><door>  EE    <serial, 8 characters>  EE    <epoch, uint32 little endian>  CRC  FC

Fields

| Field | Meaning |
| --- | --- |
| reading | A fault byte, then the value multiplied by 100 as ASCII — variable width, no leading zeros, a leading minus when negative. A channel in fault sends its fault byte alone, with no value after it. |
| EE | Field separator |
| alarm | One byte per channel: 0 none, 1 upper, 2 lower |
| door | Two identical bytes: 0x0E open, 0x0F closed |
| serial | The last 8 characters of the serial number |
| epoch | The device clock, seconds since 1 January 1970, little endian |

Which three channels depends on the build, in the same order as everywhere else: DP1, DP2, DP3, or DP1, temperature, humidity.

Parse the tail by offset, not by splitting on the separator.

Everything from the alarm bytes onwards is fixed width, so it can be measured back from the terminator: 4 epoch, separator, 8 serial, separator, 2 door, separator, 3 alarms. Only the three readings at the front are variable, and only those should be split on 0xEE. The reason is that a faulted channel’s fault byte is written over whatever the request left in the buffer, so it can hold any value — including 0xEE, which would split a field in two.

## 10.5  Indexed parameters

Four parameters carry index bytes between the identifier and the value. The value field is a fixed five characters, and a negative value carries its sign inside those five — so it has one digit fewer.

| ID | Layout of a write | Value field |
| --- | --- | --- |
| 0x6E | FF ID 11 6E <ch> <5 chars> | tenths of a Pa |
| 0x72 | FF ID 11 72 <ch> <slot> <5 chars> | tenths of a Pa |
| 0x71 | FF ID 11 71 <ch> <sign> <5 digits> | hundredths of a Pa |
| 0x5F | FF ID 11 5F <ch> <sign> <5 digits> | hundredths of a Pa |

Channel is the ASCII digit 0, 1 or 2 for DP1, DP2, DP3. Slot is 0 to 5.

0x71 and 0x5F are the odd ones out: their sign sits in a byte of its own ahead of a full five digits, where 0x6E and 0x72 keep the sign inside the five.

DP1 clamp to 300.0 Pa      FF 01 11 6E 30 30 33 30 30 30 <crc> FE
                                         0  0  3  0  0  0
  DP1 slot 3 to -2.5 Pa      FF 01 11 72 30 33 2D 30 30 32 35 <crc> FE
                                         0  3  -  0  0  2  5
  DP1 zero offset 1.25 Pa    FF 01 11 71 30 2B 30 30 31 32 35 <crc> FE
                                         0  +  0  0  1  2  5

A read takes the index bytes but no value: FF ID 10 72 <ch> <slot> CRC FE.

0x71 and 0x5F need the CUSTOMER password

Both are accepted only while customer calibration is open. The factory password unlocks calibration but not these two, and a write made under it is discarded in silence - no INVALID_PARA, no error of any kind. The stored value simply does not change, and a host that does not read back afterwards will believe it worked. Read them back; it is the only way to know.

The two do different things to the reading. The zero offset is added outright, so a write shows up immediately. The span factor is eased in a fifteenth at a time and only while the reading is above 1.5 Pa, so it trims the span without disturbing the zero - and a fresh write will not move the value at once, which is easily mistaken for a write that failed.

The per-slot offset is added to a reading whose magnitude falls in that slot. The bands are fixed:

| Slot | Applies when the reading is |
| --- | --- |
| 0 | below 50.0 Pa |
| 1 | 50.0 to 100.0 Pa |
| 2 | 100.0 to 150.0 Pa |
| 3 | 150.0 to 200.0 Pa |
| 4 | 200.0 to 250.0 Pa |
| 5 | 250.0 to 300.0 Pa |

An offset is limited to ±500.0 Pa; outside that the device sets INVALID_PARA and stores nothing.

## 10.6  Calibration  (0x30, 0x31, 0x67, 0x32, 0x33)

Calibration has to be unlocked first, with 0x37 for factory or 0x38 for customer. The window stays open for about 60 seconds and every calibration exchange restarts it.

Writing — send the TRUE value, not a correction

request   FF  ID  11  32  <5 chars, tenths>  CRC  FE

The device takes what you send as the true reading and stores the difference from what it is measuring at that moment, as measured minus reference. It also resets that channel’s recorded minimum and maximum. Sending 00250 to 0x32 therefore means “you are actually reading 25.0 right now”.

Reading — always 12 bytes

The reply carries the stored correction and nothing else, whatever the unlock state. The correction is the first five characters of the payload, zero padded, in tenths, with a minus sign in the first position when negative.

Up to V1.0.9 this reply grew with the unlock state - 39 bytes with factory calibration open and 72 with customer calibration open - because the calibration date and certificate history rode along behind the value. That history has been removed from the firmware. A host that still reads by the old lengths will wait for bytes that are never sent.

### 10.6.1  Which unit a temperature is in

This is not uniform, and it catches people out. The device stores and transmits temperature in Celsius almost everywhere; only the values a person reads off the display follow the selected unit.

| Value | Unit | Converts with 0x2E? |
| --- | --- | --- |
| Live temperature in the 0x48 frame | always Celsius | no |
| Temperature in the logs and the RAM buffer | always Celsius | no |
| Recorded minimum and maximum, 0x20 / 0x21 | the displayed unit | yes |
| Alarm setpoints, 0x09 to 0x0C | the displayed unit | yes |
| Calibration reference you WRITE to 0x32 | the displayed unit | yes |
| Calibration correction you READ from 0x32 | always Celsius | no |

So a host showing Fahrenheit has to convert the live value itself, while the minimum and maximum arrive already converted. Converting those a second time is the obvious mistake.

The last two rows are the subtle pair. The reference value you send is an absolute temperature in whatever unit the display is set to, but the correction the device derives and hands back is a difference, and a difference is held in Celsius. A difference converts by the 1.8 scale alone - the 32 degree offset belongs to absolute temperatures only. Firmware before 1.0.5 applied the absolute formula to that difference in four places, so calibrating while the display was in Fahrenheit stored a correction that was wrong by tens of degrees, and merely switching the unit altered a calibration that was already stored.

Firmware before 1.0.4 sent nothing at all for a locked calibration read — the branch built a reply and never transmitted it, leaving the host to time out. Fixed in 1.0.4.

## 10.7  Feature word  (0x4E)

read      FF  ID  10  4E  PPPP  CRC  FE
  write     FF  ID  11  4E  PPPP  BBBBB  CRC  FE

Four digits of factory password, and for a write five more holding the bit pattern in decimal. The device saves it and restarts a few seconds later, so a host has to wait for it to come back before reading anything else. On the bench it answered again about two seconds after the write.

Which parameters the instrument runs:

| Bit | Meaning | Build |
| --- | --- | --- |
| 0x0001 | DP1 channel | both |
| 0x0002 | Clock | both |
| 0x0004 | Alerts | both |
| 0x0008 | DP2 channel | DP1+DP2+DP3 only |
| 0x0010 | DP3 channel | DP1+DP2+DP3 only |
| 0x0020 | Humidity | DP1+Temp+RH only |
| 0x0040 | Temperature | DP1+Temp+RH only |

The upper four bits do not mean the same thing in both builds, so a host has to know which one it is talking to before interpreting them - the temperature-unit probe in section 10.3 settles that. Reading 0x0040 as "DP3" on a Temp+RH unit is the mistake this invites.

The clock bit is the one to be careful with. Every log refuses to write without it, so clearing it stops all logging while leaving the instrument otherwise healthy.

Logging itself is no longer switched here. It is now settled when the firmware is built, by the BUILD_REGULAR_LOG, BUILD_LOG24_LOG, BUILD_MINMAX_LOG, BUILD_MEAN24_LOG and BUILD_RAM_BUFFER macros, and a unit that should not log needs a different build rather than a different setting.

## 10.8  Reading the logs

Five logs, all written only while the clock has been set AND is trusted - checked when each record is written, not merely at start-up. A clock that reads plausibly is not enough: it has to have been set deliberately, which is what raises the flag the writers look for. After a first-time initialisation the instrument will run happily and record nothing until its clock is set.

| Log | Command | Reply frame | Holds |
| --- | --- | --- | --- |
| Regular log | 0x49 RDLG_DT | 70 bytes | 60000 records |
| 24 hour ring | 0x47 FLASH24_IND | 72 bytes | 1440 records |
| 15 day min/max/mean | 0x53 MINMAXMEAN_IND | 25 bytes | 15 days per channel |
| 24 hourly means | 0x54 MEAN_HR | 13 bytes | 24 values per channel |
| RAM buffer | 0x45 RAM_ALL | 1507 bytes | last 30 readings |
| RAM buffer, by slot | 0x46 RAM_IND | 71 bytes | a range of the 30 |

Four of the five are STREAMED

The request arms a transfer and the device then pushes one frame per pass of its main loop until the count is used up. So a read is one request followed by however many frames arrive, not a request and a single reply. RAM_ALL is the exception: one reply carrying all thirty records.

While a transfer is running the device serves no other command, and it releases by itself when the count reaches zero.

Requests

0x49  FF ID 10 49  dd mm yy hh mi ss  dd mm yy hh mi ss  CRC FE
                     \___ start ____/  \____ end ____/   raw bytes

  0x47  FF ID 10 47  SSSS EEEE  CRC FE      start and END slot, 4 ASCII digits each
  0x53  FF ID 10 53  c NN       CRC FE      channel 0-2, count as 2 ASCII digits
  0x54  FF ID 10 54  c          CRC FE      channel 0-2; the count is fixed at 24
  0x46  FF ID 10 46  ss ee      CRC FE      start and END slot, raw bytes

0x47 and 0x46 take a RANGE, not a count

Both indices are slot numbers and both ends are included, so start 0 end 9 returns ten records. The walk runs forwards and wraps round the end of the store, so a start above the end is a window that spans the wrap rather than an error. Records therefore arrive oldest first within the window.

An index outside the store - past 1439 on the ring, past 29 in the RAM buffer, or any non-digit in the four ASCII characters 0x47 expects - is answered with INVALID_PARA and no data. Earlier firmware checked neither: four ASCII digits reach 9999 against a 1440 slot ring, and on 0x46 the index became a byte offset into a 2000 byte buffer, so a stray value streamed whatever memory followed it.

Up to and including V1.0.8 each of these parsed one of its two indices and then discarded it, using the other as a COUNT: 0x47 forced the start to the newest record and read the second field as how many to send, and 0x46 did the same the other way about. A host that sent start 0 appeared to work and was not doing what it looked like. There is no "newest N records" command - work the range out from the ring index, 0x4C, which is where the NEXT record will go.

A regular-log transfer leads with a short frame giving how many records fall in the window, as a uint32 - not the ASCII the ordinary replies use.

That count includes BOTH ends of the window. Up to V1.0.8 it was the difference between the two indices and so was one short, and the record it left out was always the newest one in the range.

An empty log is answered with that same frame carrying zero. Earlier firmware sent nothing at all, so a host asking a fresh unit could not tell "no records" from "no reply" and simply waited out its timeout.

A window whose END precedes its START is refused with INVALID_PARA

The two dates used to be compared only against what the log held, never against each other. A reversed window passed both of those tests, the two index searches then came back in descending order, and the span arithmetic read that as a wrap round the ring - so the device streamed very nearly all 60000 records, about an hour and fifty minutes, serving no other command for any of it. The host cannot call a transfer off once it is running.

Where the record sits in the frame

This differs between them, and a wrong offset still parses and yields plausible nonsense rather than failing:

| Command | Record at | Why |
| --- | --- | --- |
| 0x49 regular log | byte 5 | nothing between the header and the record |
| 0x47 24 hour ring | byte 7 | a 2 byte index first, big endian |
| 0x46 RAM buffer | byte 6 | a 1 byte index first |
| 0x45 RAM_ALL | byte 5 + n*50 | thirty records end to end |
| 0x53 min/max/mean | byte 7 | 16 byte record |
| 0x54 hourly mean | byte 7 | a single float |

The 50 byte reading record

The regular log, the 24 hour ring and the RAM buffer all share one layout:

[0..3]   epoch seconds, uint32 little endian
  [4]      device id
  [5]      log type
  [6]      user id
  [7..8]   password
  [9]      sensor fault flags, same bits as the status byte
  [10..13] DP1                        float
  [14..17] DP2 or temperature         float
  [18..21] DP3 or humidity            float
  [22..29] DP1 minimum, maximum       float, float
  [30..37] DP2 / temperature min, max
  [38..45] DP3 / humidity min, max
  [46..48] the three alarm states, 0 none 1 high 2 low

Temperature in a record is always Celsius, whatever the display was set to. Bit 7 of [47] records the unit that was on the display, but the value is not converted.

The 16 byte day record  (0x53)

[0..3]   epoch of the day
  [4..7]   minimum     float
  [8..11]  maximum     float
  [12..15] mean        float

Slots that have never been written read back as erased flash, which decodes to timestamps in the 1970s and values with twenty digits. A host should reject anything whose epoch predates the product rather than display it.

# 11.  Things that will catch a host out

The two index reads, 0x47 and 0x46, take a start and an END slot, both ends included, and refuse an index outside the store with INVALID_PARA. Firmware up to V1.0.8 read the second field as a count instead, so a host written against that behaviour will ask for the wrong records rather than fail.

Ask 0x49 for a window the right way round. An end earlier than the start is refused now, but older firmware answered it by streaming the whole log.

The duration sent with 0x3F is stored and then ignored. The countdown that would have stopped broadcasting when it expired is not compiled, so broadcasting runs until 0x3E stops it. Send a duration by all means, but do not rely on it.

0x69 is in SECONDS, not minutes — the device counts it down once a second. It was documented as minutes until V1.2.2.

A write reply confirms the frame, not the value. A value outside the parameter’s range is discarded without any error, so read it back.

Check the parameter identifier in byte 4 of every reply. If a reply is ever lost, the next one arrives against the following request and a value shows up under the wrong parameter’s name — plausible numbers, silently wrong.

Changing the device address (0x1A) needs the same value written twice within about ten seconds. The first write only arms it.

Broadcast, address 0, is not usable for reads. Every device on the bus answers and the replies collide. The reply also carries the responding device’s own address, never 0.

Disabling the UART (0x70 set to 1) is one-way over the link: once it is off the only frame the device will still act on is a write of 0x70 itself.

The log interval (0x19) is only loaded from storage on a unit that has logging enabled. On any other unit a read returns whatever is in RAM, which is 0 — and 0 cannot be written back, because the accepted range starts at 1.

Reading a parameter that belongs to the other build returns INVALID_PARA, not silence. Silence means the request was lost; see section 7.

Every value is parsed into a signed 16 bit integer. Do not send anything outside −32768 to 32767.

The DP zero offset (0x71) and span factor (0x5F) are accepted only under the customer password. Under the factory password the write is discarded silently, with no error, so always read them back.

A log transfer takes over the link. While one is running the device serves no other command, so a host must read the stream out before asking for anything else.

Records cannot be found by scanning for the 0xFC terminator. The floats inside them contain that byte regularly, so frames have to be taken by their fixed length.

Temperature is not carried in one consistent unit. The live frame and the logs are always Celsius; the recorded extremes and the alarm setpoints follow the displayed unit. See 10.5.1 before writing any conversion.

A recorded DP extreme that lies outside the DP clamp cannot be superseded - a minimum only moves down and a maximum only moves up, and the clamp stops any reading reaching it. Firmware 1.0.5 restarts such an extreme from the present reading; before that, narrowing the clamp stranded the pair at values the device could no longer produce.

# 12.  Defaults on a fresh device

| Setting | Default |
| --- | --- |
| Device address | 1 |
| Baud rate | 57600, code 8 |
| Customer password | 100 |
| Factory password | 1000 |
| Factory parameter-set password | 1234 |
| Log interval | 1 minute |
| DP reading clamp | 300.0 Pa |
| Per-slot DP offsets | all zero |
| Serial number | sixteen ASCII zeros |
| Clock epoch | 1 January 1970 |

# 13.  This build

| Item | Value |
| --- | --- |
| Firmware version | 1.2.2 (0x34 returns 122) |
| Channel layout | DP1 + Temperature + Humidity |
| Pressure sensor | XGZP6891D |
| Data flash | AT45DB321D |
| Differential pressure rating | ±500 Pa |
| Microcontroller | MM32F0141C6P, 64 KB flash, 8 KB RAM |

Generated from the firmware sources. Parameter access and build applicability were read out of the dispatch tables in main.c rather than written by hand.